**Date- 16.11.2023**

**Tender Reference No.:- GEM/2023/B/3142973**

**Technical Corrigendum**

**Name of Equipment: General Ventilator**

|  |  |  |  |
| --- | --- | --- | --- |
| **Sr. No.** | **Reference Of Tender Clause** | **Tender Clause** | **Revised Clause** |
| **1.** | Technical Specification of General Ventilator | Certification Number and date : TC561818000000014F and Dared 31-10.2018 | Deleted |
| **2.** | Technical Specification of General Ventilator | Volume Support | Volume Support:-Optional |
| **3.** | Technical Specification of General Ventilator | The Unit should be External compressor based on same make for precise gas delivery. | Compressor/Turbine BasedNote:-For compressor based, compressor should be of same make (OEM) |
| **4.** | Technical Specification of General Ventilator | The Ventilator model should be EU CE certified. | 1- Certifications for the product:- EU CE (From 4 digit notified body)/USFDA/BIS2- Manufacturer Should Be ISO 134853-CDSCO Registered/Certificate |
| **5.** | Technical Specification of General Ventilator | Inspiratory pressure, cm H2O: up to 150 or more | Inspiratory pressure, cm H2O should be up to 100 or more |
| **6.** | Technical Specification of General Ventilator | Facility for double lung ventilation: Yes | Facility for double lung ventilation :-Optional |
| **7.** | Technical Specification of General Ventilator | USB/Memory card: Yes | Deleted |
| **8.** | Technical Specification of Advance Ventilator | Facility to Report (vent alarms and patient status): Yes - | Facility to Report (vent alarms and patient status): Optional |
| **Note:-**Flow sensor should be reusable Supplier has to replace the Flow sensor during warranty Period Free of cost. Warranty of flow sensor should be considered covered as per ventilator warranty period, 2 nos. of flow sensors should be supplied along with each equipment. |